## Slide 1
CONCERT Project Website (https://concert-h2020.eu)
CONFIDENCE Project Website (https://portal.iket.kit.edu/CONFIDENCE//)
NERIS Platform Website (https://www.eu-neris.net/projects.html)
Evaluation of apps for dose-rate measurements using the camera sensor of the smartphone
Alessia Mafodda, Clemens Woda
Institute of Radiation Medicine, Helmholtz Zentrum München , Munich, Germany
alessia.mafodda@helmholtz-muenchen.de
2 - 5 December 2019
Lindner Hotel Gallery Central
Bratislava, Slovak Republic
CONFIDENCE Dissemination Workshop
“Coping with uncertainties for improved modelling and decision making in nuclear emergencies”
Introduction
Objectives
Materials and methods
Conclusion
Dose-rate response
Some applications available on the market turn a smartphone into a dose-rate meter using the Complementary Metal Oxide Semiconductor (CMOS) sensor of the camera as a radiation detector. The working principle behind is that CMOS sensor is sensitive to visible light but can detect ionizing radiation when shielded with adhesive black tape1).
After covering the camera, the user may run a measurement that consists of a video record where every interaction of radiation (photons) with the chip is observed as an intense signal in a certain pixel (bright spot). The dedicated software analyses each video frame by counting the number of spots, and the sum of detected particles in selected time periods is given as number of counts.
The characterization of two applications (“RadioactivityCounter” and “GammaPix”) was carried out in order to assess the reliability of such “citizen measurements”.
Tests on:
Dose-rate response of all the models for dose-rates between 2 to 1000 μGy h-1
Background assessment (when possible)
Energy dependence - tests on front/rear cameras
Angular dependence for different impact angles (0° = back camera facing the source and 180°= back camera away from the source)
“GammaPix”
For all the tested models the App developer provides an
approximate calibration.
Background is established in three-step procedure
and it‘s not adjustable.
A “warning message” is displayed and stored at
the end of every measurement, representing
a fast and clear way to communicate people
radiation levels and effects
Reading Unclear / All Clear / Radiation Detected
More efforts were put into evaluating the “RadioactivityCounter” app, since the lack of calibration of “GammaPix” resulted in underestimating the nominal reference values on most of the phones by at least an order of magnitude.
The „noise level“ (n≶10) influences the dose-rate responses and assessing it in a shielded environment does not improve the performances at lower dose-rates
Software algorithm for image corrections leads to a reduced sensitivity of the sensor for ionizing radiation
Smartphones apps can potentially identify regions of high radioactive contamination for measurement times of 10 minutes, whereas reliable measurements at natural environmental radioactivity level are difficult to assess
Wp2
“RadioactivityCounter”
Calibration not available for most of the models, but adjustable
“Set the noise” procedure for assessing the
n level that optimizes the ratio of sensitivity
for radiation to sensor noise
Allows to log data that can be downloaded
as .csv files containing date, time, count rate
and, if calibrated, also the dose-rate, for
each minute as well as averaged
over the measurement period.
14 different mobile phones from 2017
Irradiations performed at the SSDL of the Helmholtz Zentrum München with 137Cs and X-ray sources
Energy-dependence
Chart: ASUS Zenfone 3 and Lenovo K6 comparison
| 24 | 33 | 48 | 65 | 100 | 118 | 165 | 208 | 250 | 660
ASUS Zenfone 3 | 0.1827 | 1.6445 | 7.6462 | 10.6508 | 6.4492 | 5.4869 | 2.8411 | 1.8192 | 1.5789 | 1
Lenovo K6 | 0.75868 | 3.5818 | 9.3762 | 10.4784 | 7.5639 | 6.0611 | 3.2095 | 2.2071 | 1.6301 | 1
Chart: iPhone 6S - front and rear camera
| 24 | 48 | 65 | 100 | 118 | 208 | 250 | 660
rear camera facing source | 1.5501 | 15.9278 | 17.9177 | 13.7395 | 10.7816 | 2.8424 | 2.2513 | 1
front camera facing source | 1.1273 | 14.9123 | 15.4894 | 1.5606 | 1
front camera facing user | 0.045665 | 8.3137 | 12.2641 | 1.3648 | 1
Background assessment
A low noise level (n<10) often characterizes phones in the “high price” range (over 200 €) for which a linear trend in dose-rate responses with proportionality starting from 5 μGy h-1 is observed for integration time of 10 minutes (Fig.1 - left).
One hour-measurements slightly improve the linearity down to 2 μGy h-1 but counts distribution is very skewed (Fig.1 – right).
Phones with high noise level (n>10) show linear trend in dose-rate responses with proportionality starting from 50 μGy h-1 for integration time of 10 minutes (Fig.2 – left). One hour-measurements don’t improve the “blindness” below that threshold, but counts are more normally distributed (Fig.2 - right).
Assessing the background in a shielded environment (20 nSv/h) leads to an increase in the number of counts (Fig.4)but not to an improvement in detecting dose-rates lower than 10 μGy h-1.
Despite the different models and the quality of the camera sensors, all models showed a strong over-response of up to a factor of 18 for a photon energy of around 60 keV. A comparison was carried out between two phones (ASUS Zenfone 3 and Lenovo K6) mounting the same sensor type (Fig.3 - left). The relative response of the back and front camera was compared for three models (Fig.3-right)
Chart: iPhone 6S "RadioactivityCounter"
| 2 | 5 | 10 | 20 | 50 | 100 | 500 | 1000
10 minutes based | 4.562 | 4.7513 | 10.1022 | 22.7907 | 54.7407 | 111.24 | 546.3311 | 1052.815
 | 1 | 1000
60 minutes based | 1.7322 | 5.5293 | 12.2583
Chart: 
Fig.1
Chart: 
Chart: HUAWEI P8 lite "RadioactivityCounter"
| 2 | 5 | 10 | 20 | 50 | 100 | 500 | 1000
10 minutes based | 13.8182 | 13.7273 | 20.625 | 29.5 | 44.625 | 72.875 | 362.6667 | 659.625
60 minutes based | 14.3684 | 15.6429 | 16.75 | 23.4821 | 37.3158
 | 1 | 1000
Fig.2
Fig.3
Chart: iPhone 6S
| 2 | 5 | 10 | 20 | 50 | 100 | 500 | 1000
unshielded | 4.562 | 4.7513 | 10.1022 | 22.7907 | 54.7407 | 111.24 | 546.3311 | 1052.815
shielded | 15.9867 | 23.9673 | 25.8589 | 25.9692 | 112.9167 | 154.124
Fig.4